General Monthly Summary

Overall data completeness for December was 99.2%. The T17 trap was replaced November 28, which caused an unusual baseline disturbance, distorting the methylcyclopentane and 2,4-dimethylpentane peaks. This disturbance persisted through December 8. A calibration was performed December 1. The new PLOT column RF differed by 16.8% from the previous RF. The new BP column RF differed by 13.6% from the previous RF. Four ambient hours were nulled because data was missing. One ambient hour was nulled because of aberrant chromatography.

Data Qualification Summary

Acetylene and p-diethylbenzene were qualified with QX (Does not Meet QC Criteria) for the entire month due to failing CVS recoveries. N-dodecane, 1-hexene and isoprene were qualified with LJ for the entire month due to poor precision for the heaviest compounds in the BP column. 2-methyl-1-pentene, alpha-pinene, and beta-pinene were qualified with LJ because there was no reference gas available. 2-methylpentane was qualified LJ for the entire month because of persistent overlap from a contaminant peak. m-Diethylbenzene and p-diethylbenzene were qualified with flag LJ for the month due to uncertainty at low concentrations.

Data Null Summary

The following ambient hours had all analytes nulled.  Four hours were nulled because no data was collected for an unknown reason (AM). One hour was nulled because of significant RT shifts and baseline noise (DA).

Methylcyclopentane and 2,4-dimethylpentane were nulled with flag DA between December 1 and December 8 because of overlap with a baseline disturbance.

QA Summary

Blank Response

Ethylene exceeded its MDL for the entire month and was above 0.5 ppbC in eight blanks between December 1 and December 6. Ethane exceeded its MDL periodically throughout the month (11 exceedances) until December 17, after which it recorded no further exceedances. Ethane also exceeded 0.5 ppbC once during the month (December 14). Acetylene exceeded its MDL in almost every blank between December 1 and December 17 (13 exceedances), after which it recorded no further exceedances.  All compounds with MDL exceedances were qualified with flag LB forward and backward to the nearest passing blanks. All compounds above 0.5 ppbC were nulled with flag AS forward and backward to the nearest passing blank.

Calibration Verification Standard (CVS) Recoveries

Acetylene and p-diethylbenzene failed all month due to poor tank retention and were qualified with flag QX (Does not meet QC criteria). After December 23, 1,2,4-tri-m-benzene began consistently recovering below 70% and was flagged accordingly. Precision checks passed for the entire month (<25% RPD) for all compounds, but were elevated for heavier VOCs on the BP column, especially p-diethylbenzene. Duplicate CVS runs for these values were consistently elevated relative to the initial runs, consistent with observations at other sites.  The target for propane and toluene was 1.36 ppbC. The standard deviation for propane CVS runs was 0.03 ppbC and for toluene was 0.04 ppbC.

Retention Time Standard (RTS) Recoveries

Acetylene failed all month due to poor tank retention.  The RTS showed decent recovery throughout the month. The standard deviation for propane was 0.03 ppbC and for toluene was 0.04 ppbC.

Laboratory Control Standard (LCS) Recoveries

Acetylene was below 70% recovery in all LCS runs for the month. The standard deviation for propane was 0.11 ppbC and for toluene was 0.09 ppbC.

Calibrations

A calibration was performed on December 1. The new RF values were 4872 for the PLOT column and 5025 for the BP column. The previous RF values were 5691 for the PLOT column and 5708 for the BP column. The relative percent difference between these RFs was 16.8% for the PLOT column and 13.6% for the BP column.

Method Detection Limit (MDL) Studies

No MDL studies were performed this month.

Proficiency Testing (PT)

No PT analysis was performed this month.

Ambient Air Spikes

One ambient spike was performed on December 20.

Nonconformances

None recorded.

Trends

QC recoveries for propane and toluene were consistent and within acceptable bounds for the entire month.