## Validator Notes
Site | EQ - Utah Technical Center
Month/Year | Month 2025
Validator | Aengstrom
Validation Prep: |  |  |  | Notes:
Verify all Daily-ZIP folders are downloaded into U:\PLAN\AMC\GC-Agilent+Markes\Data\ORIGINAL\{YYYY}\{YYYYMM}
Verify Operator Reprocessed zip folders are downloaded into U:\PLAN\AMC\GC-Agilent+Markes\Data\ORIGINAL\{YYYY}\{YYYYMM}
Verify all data has been extracted to C:\Users\{Username}\Documents\AutoGC Data\{SiteID}\{YYYY}\{YYYYMM}
Ensure dates, times, and site names/numbers in files match up and are continuous. document any missing files on the Validation Notes.
Verify all Original TX1 files loaded into MAX
Check CVS, Blanks, and LCS Reports and notate failures
Download Crosstab into MDVR folder pre changes
Review Oplog:
View and Print the Operator Log Book (Desktop), Op Log backup (D:\EZ\HW-meth), and Operator Checklist (Desktop) from the HW-GC computer with 72 hours on either side of the month. Review and highlight major events in the operator log book such as optimizations, cylinder changes, major events (calibrations), major malfunctions (power failures, trap, oven door failures), etc. Highlight hours for review during or around major events as necessary per operator log book.
Stock Cylinder Changes - Verify in MAX
Correct PE Audit Samples
Correct MDL Samples - Create Spreadsheet
Annual PM
3pt Curve - Calculated RPD
Review Quicklooks:
Dates (1-31) Hours (A-X)
Invalidate missing hours
Over-ranges - If valid: remove flag, recal & store, update excel
Invalidate data collection errors
Pick hours to view (1st/last month, surrounding invalidations)
Highlight at least 3 hours to review daily on each column’s hourly grid page. Ideally one hour will be reviewed for each of the following time span; Morning Commute (06-09), Afternoon (11-15); Night (0-4, or16-23). Try not to exceede 6hr.
Review Chromatography: |  |  |  | CVS Std | BR | LCS Std | DF | RTS Std | BR | 1) RSD (Plot) =
All QC samples have been viewed
| All highlighted hours have been viewed | [8 empty] | 1) RSD (BP) =
| If reprocessed, viewed reprocessed QC/Ambient |  |  | Starting RF |  | Ending RF |  | ReCal Date(s)
| View hours before & after reprocessed data |  |  | PLOT | BP1 | PLOT | BP1 | ReCal Date(s) |  | 2) RSD (Plot) =
View hours before & after invalidations
| Invalidate TMNHC/TC if invalidating 8+ compounds | [8 empty] | 2) RSD (BP) =
| >336 hour, 14-days, of data loss; update Reprocess Plan |  |  | EH =
Send necessary emails and save copy in the MDVR folder for general issues, 336-hr data losses, and ambient over-ranges.
Validator Notes
Site | EQ - Utah Technical Center
Month/Year | Month 2025
Validator | Aengstrom
Review Reprocess Plan: |  |  |  | Notes:
Number of tx1s matches Reprocess Data
Reprocessed Data matches Reprocess Plan
All QC failures listed as Invalidations or Qualified
All hourly invalidations recorded
Validation:
Load and Decode new TX1 files in MAX
Final QC report reviewed for addition invalidations
Review new Quicklooks & compare with Reprocess Plan
Flag invalid hours, components
Validate & Lock data in MAX
Create AQS file (targets!) ( Include QA/QC, Include X Run, Do not Include TNMTC (43000) or Show amount if nulled data)
Create AQS Blank file
Download Crosstab into MDVR folder post changes
Email AQS File to Data uploader (Bart)
Calculate and store recoveries below
Zip data folder (including raw, rst, tx0, mtd, seq)
Move Xyymmdd.zip to U:\PLAN\AMC\GC-Agilent+Markes\Data\FINAL
File Archive and AQS Upload
| AQS File Date | AQS File Name (MMMYYYYXX) | Total # of Records
| Applied Validation Lock | Archive File Name (HWyymmdd.zip) | Date File Archived
Data Completeness Report from Orset MAX
| Date data recovery | Total Valid Ambient Hours | Total Null Flagged QC | Total Invalid Hours
| Hours in Time Period | Valid Ambient Recovery (%) | Totl Valid Ambient Plus QC | % Data Return adjusted for QC
|  | #DIV/0! | 0 | #DIV/0!
*Report is based on the sample name in the tx0 or tx1 file.
Qulaifier Flag, MD, was applied to all paramaters and hours that measured below the MDL; Qulaifier Flag, ND, was applied to all paramaters and hours that where not detected; Qulaifier Flag, EH, was applied to Isoprene, 2-METHYL-1-PENTENE, 1-Hexene, and n-Dodecane - due to questions about the accuracy of the measurments at the end of the column; Qualif Flag, QX, was applied to ACETYLENE for whole month due to low CVS recovery; Qualif Flag, QX, was applied to BETA-PINENE from 5-1 to 5-16 due to high blank value;

## Reference Info
| Site | EQ - Utah Technical Center |  | Validator |  | Aengstrom
| Month/Year | Month 2025
*See Validation SOP for calculations.
Standards and References: |  |  |  | PLOT |  |  |  | BP |  |  |  | Calibrant Target
Standard | Canister Lot# | Blend Ratio | Start Date | Std Dev (ppbC) | Average (ppbC) | % Bias | % Error | Std Ave (ppbC) | Average (ppbC) | % Bias | % Error | PLOT (ppbC) | BP (ppbC) | Notes:
RTS | CC731198-0126 | 0.0013 | 2024-03-08 | 0.091125 | 2.0121 | 148.8% | 6.74% | 0.076813 | 2.0227 | 1.4406 | 5.471% | 1.352 | 1.404
CVS | CC524930-0626 | 0.00188 | 2025-08-15 08:21 | 0.069578 | 5.8172 | 303.4% | 3.628% | 0.064025 | 5.8613 | 305.1% | 3.332% | 1.9176 | 1.9214
CVS | CC524930-0626 | 0.00161 | 2025-09-08 17:00 | 0.061229 | 1.1159 | 67.95% | 3.728% | 0.033735 | 1.494 | 90.8% | 2.05% | 1.6422 | 1.6454
LCS | CC177206-1226 | 0.005 | 2025-06-03 12:25 | 0.066764 | 1.0806 | 20.01% | 1.236% | 0.038906 | 1.495 | 27.99% | 0.7284% | 5.4 | 5.341
| RF |  | RSD |  |  | EH (Calibration Range)
Calibration | PLOT | BP | PLOT | BP |  | PLOT (ppbC) | BP (ppbC) | Reference File: |  |  | Notes:
2025-09-03 | 6104 | 5828 | 0% | 0% |  | 44.63 | 44.63 | 07242025 MMSD-MPV-23046_PPMC
2025-09-10 | 5925 | 5623 | 3.021% | 3.646% |  | 44.63 | 44.63

## QC Review
QC (CVS, CVS Precision, LCS, RTS, BLANK, Audit) Review
Site | EQ - Utah Technical Center | Validator | Aengstrom
Month/Year | Month 2025
QC Type | Date/Time | Filename | Plot Column Notes | BP Notes | Actions
RTS
LCS
CVS
CVS Precision
Audit
Blank
Exp. Or Test Sample
Qulaifier Flag, MD, was applied to all paramaters and hours that measured below the MDL; Qulaifier Flag, ND, was applied to all paramaters and hours that where not detected; Qulaifier Flag, EH, was applied to Isoprene, 2-METHYL-1-PENTENE, 1-Hexene, and n-Dodecane - due to questions about the accuracy of the measurments at the end of the column; Qualif Flag, QX, was applied to ACETYLENE for whole month due to low CVS recovery; Qualif Flag, QX, was applied to BETA-PINENE from 5-1 to 5-16 due to high blank value;

## Reprocessed Data
Site | EQ - Utah Technical Center
Month/Year | Month 2025
Validator | Aengstrom
|  | Batch |  | single file rpo | Data Range |  |  |  |  | Date
Column | Corrections Made | Sequence | method | data file | Date | Hour | - | Date | Hour | RPO
|  | Batch |  | single file rpo | Data Range |  |  |  |  | Date
Column | Corrections Made | Sequence | method | data file | Date | Hour | - | Date | Hour | RPO
|  | Batch |  | single file rpo | Data Range |  |  |  |  | Date
Column | Corrections Made | Sequence | method | data file | Date | Hour | - | Date | Hour | RPO
Qulaifier Flag, MD, was applied to all paramaters and hours that measured below the MDL; Qulaifier Flag, ND, was applied to all paramaters and hours that where not detected; Qulaifier Flag, EH, was applied to Isoprene, 2-METHYL-1-PENTENE, 1-Hexene, and n-Dodecane - due to questions about the accuracy of the measurments at the end of the column; Qualif Flag, QX, was applied to ACETYLENE for whole month due to low CVS recovery; Qualif Flag, QX, was applied to BETA-PINENE from 5-1 to 5-16 due to high blank value;

## Reprocess Plan
Reprocess Plan | [9 empty] | Site: |  | EQ - Utah Technical Center | [5 empty] | Month/Year: |  |  |  | Month 2025 | [24 empty] | Pg. | 1
| 1 | [23 empty] | 2 | [24 empty] | CVS
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 |  | Blank
| A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X | A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X |  | LCS
Checked PLOT | [49 empty] | RTS
RPO PLOT | [49 empty] | Experimental/ Test
Checked BP | [49 empty] | Missing Data
RPO BP | [49 empty] | Invalidate
Invalid PLOT | [49 empty] | Check this Hour
Invalid BP | [48 empty] | RP | Reprocessed file
Op Log | [48 empty] | OK | Hour reviewed and correct
Notes: | [49 empty] | Multipoint
| 1 0 A: 3 | 4 | [24 empty] | Detection Limit
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | SB | See Below
| A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X | A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X
Checked PLOT
RPO PLOT
Checked BP
RPO BP
Invalid PLOT
Invalid BP
Op Log
Notes:
| 3 0 A: 5 | 6
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23
| A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X | A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X
Checked PLOT
RPO PLOT
Checked BP
RPO BP
Invalid PLOT
Invalid BP
Op Log
Notes:
| 5 0 A: 7 | 8
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23
| A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X | A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X
Checked PLOT
RPO PLOT
Checked BP
RPO BP
Invalid PLOT
Invalid BP
Op Log
Notes:
Reprocess Plan | 7 9 J: Site:; 7 11 L: EQ - Utah Technical Center; 7 17 R: Month/Year:; 7 21 V: Month 2025 | [22 empty] | Pg. | 2
| 7 0 A: 9 | 10 | [24 empty] | CVS
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 |  | Blank
| A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X | A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X |  | LCS
Checked PLOT | [49 empty] | RTS
RPO PLOT | [49 empty] | Experimental/ Test
Checked BP | [49 empty] | Missing Data
RPO BP | [49 empty] | Invalidate
Invalid PLOT | [49 empty] | Check this Hour
Invalid BP | [48 empty] | RP | Reprocessed file
Op Log | [48 empty] | OK | Hour reviewed and correct
Notes: | [49 empty] | Multipoint
| 9 0 A: 11 | 12 | [24 empty] | Detection Limit
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | SB | See Below
| A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X | A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X
Checked PLOT
RPO PLOT
Checked BP
RPO BP
Invalid PLOT
Invalid BP
Op Log
Notes:
| 11 0 A: 13 | 14
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23
| A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X | A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X
Checked PLOT
RPO PLOT
Checked BP
RPO BP
Invalid PLOT
Invalid BP
Op Log
Notes:
| 13 0 A: 15 | 16
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23
| A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X | A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X
Checked PLOT
RPO PLOT
Checked BP
RPO BP
Invalid PLOT
Invalid BP
Op Log
Notes:
Reprocess Plan | 15 9 J: Site:; 15 11 L: EQ - Utah Technical Center; 15 17 R: Month/Year:; 15 21 V: Month 2025 | [22 empty] | Pg. | 3
| 15 0 A: 17 | 18 | [24 empty] | CVS
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 |  | Blank
| A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X | A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X |  | LCS
Checked PLOT | [49 empty] | RTS
RPO PLOT | [49 empty] | Experimental/ Test
Checked BP | [49 empty] | Missing Data
RPO BP | [49 empty] | Invalidate
Invalid PLOT | [49 empty] | Check this Hour
Invalid BP | [48 empty] | RP | Reprocessed file
Op Log | [48 empty] | OK | Hour reviewed and correct
Notes: | [49 empty] | Multipoint
| 17 0 A: 19 | 20 | [24 empty] | Detection Limit
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | SB | See Below
| A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X | A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X
Checked PLOT
RPO PLOT
Checked BP
RPO BP
Invalid PLOT
Invalid BP
Op Log
Notes:
| 19 0 A: 21 | 22
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23
| A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X | A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X
Checked PLOT
RPO PLOT
Checked BP
RPO BP
Invalid PLOT
Invalid BP
Op Log
Notes:
| 21 0 A: 23 | 24
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23
| A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X | A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X
Checked PLOT
RPO PLOT
Checked BP
RPO BP
Invalid PLOT
Invalid BP
Op Log
Notes:
Reprocess Plan | 23 9 J: Site:; 23 11 L: EQ - Utah Technical Center; 23 17 R: Month/Year:; 23 21 V: Month 2025 | [22 empty] | Pg. | 4
| 23 0 A: 25 | 26 | [24 empty] | CVS
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 |  | Blank
| A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X | A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X |  | LCS
Checked PLOT | [49 empty] | RTS
RPO PLOT | [49 empty] | Experimental/ Test
Checked BP | [49 empty] | Missing Data
RPO BP | [49 empty] | Invalidate
Invalid PLOT | [49 empty] | Check this Hour
Invalid BP | [48 empty] | RP | Reprocessed file
Op Log | [48 empty] | OK | Hour reviewed and correct
Notes: | [49 empty] | Multipoint
| 25 0 A: 27 | 28 | [24 empty] | Detection Limit
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | SB | See Below
| A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X | A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X
Checked PLOT
RPO PLOT
Checked BP
RPO BP
Invalid PLOT
Invalid BP
Op Log
Notes:
| 27 0 A: 29 | 30
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23
| A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X | A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X
Checked PLOT
RPO PLOT
Checked BP
RPO BP
Invalid PLOT
Invalid BP
Op Log
Notes:
| 29 0 A: 31 | Month Notes
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23
| A | B | C | D | E | F | G | H | I | J | K | L | M | N | O | P | Q | R | S | T | U | V | W | X
Checked PLOT
RPO PLOT
Checked BP
RPO BP
Invalid PLOT
Invalid BP
Op Log
Notes:

## QUALIFIERS_NULL
QUALIFIERS / NULL
Site | EQ - Utah Technical Center
Month/Year | Month 2025
Validator | Aengstrom
Null | [11 empty] | Common Null Flags |  | Null
[5 empty] | DATE/TIME RANGE | [6 empty] | Code | Description |  |  |  | DATE/TIME RANGE
Paramater (s) |  | COMPOUND(S) or WHOLE HOUR(S) - REASON |  | CODE | Date | Hour | - | Date | Hour | Justification | Hours | AF | Scheduled but not Collected (Lost) | Paramater (s) | COMPOUND(S) or WHOLE HOUR(S) - REASON | CODE | Date | Hour | - | Date | Hour | Justification
[6 empty] | 0 | - |  | 23 |  | 0 | AG | Sample Time Out of Limits |  |  |  |  | 0 | - |  | 23
[6 empty] | 0 | - |  | 23 |  | 0 | AH | Sample Flow Rate out of Limits |  |  |  |  | 0 | - |  | 23
[6 empty] | 0 | - |  | 23 |  | 0 | AI | Insufficient Data (Invalid Sample Time) |  |  |  |  | 0 | - |  | 23
[6 empty] | 0 | - |  | 23 |  | 0 | AM | Misc. Void |  |  |  |  | 0 | - |  | 23
[6 empty] | 0 | - |  | 23 |  | 0 | AN | Machine Malfunction |  |  |  |  | 0 | - |  | 23
[6 empty] | 0 | - |  | 23 |  | 0 | AQ | Collection Error |  |  |  |  | 0 | - |  | 23
[6 empty] | 0 | - |  | 23 |  | 0 | AS | Poor QA results |  |  |  |  | 0 | - |  | 23
[6 empty] | 0 | - |  | 23 |  | 0 | AT | Calibration |  |  |  |  | 0 | - |  | 23
[6 empty] | 0 | - |  | 23 |  | 0 | AV | Power Failure |  |  |  |  | 0 | - |  | 23
[6 empty] | 0 | - |  | 23 |  | 0 | AX | Precision Check |  |  |  |  | 0 | - |  | 23
[6 empty] | 0 | - |  | 23 |  | 0 | AY | QC Control Points, Zero/Span |  |  |  |  | 0 | - |  | 23
[6 empty] | 0 | - |  | 23 |  | 0 | AZ | QC Audit (Audit Sample) |  |  |  |  | 0 | - |  | 23
[11 empty] | 0 | BA | Maintenances / Routine Repairs
[11 empty] | 0 | BH | Interference, Co-elution, Misid
[11 empty] | 0 | BJ | Operator Error
[11 empty] | 0 | BK | Site Computer, Data Logger Down
[12 empty] | DA | Aberrant Data, Corrupt Files, Spikes, Shifts
QUALIFIERS | [11 empty] | DL | Detection Limit | QUALIFIERS
[5 empty] | DATE/TIME RANGE | [6 empty] | TC | Component Check & RTS (RTS) |  |  |  | DATE/TIME RANGE
Paramater (s) |  | COMPOUND(S) or WHOLE HOUR(S) - REASON |  | CODE | Date | Hour | - | Date | Hour | Justification | Hours | XX | Experimental Data | Paramater (s) | COMPOUND(S) or WHOLE HOUR(S) - REASON | CODE | Date | Hour | - | Date | Hour | Justification
[6 empty] | 0 | - |  | 23 |  | 0 | [6 empty] | 0 | - |  | 23
[6 empty] | 0 | - |  | 23 |  | 0 | [6 empty] | 0 | - |  | 23
[6 empty] | 0 | - |  | 23 |  | 0 | [6 empty] | 0 | - |  | 23
[6 empty] | 0 | - |  | 23 |  | 0 | [6 empty] | 0 | - |  | 23
[6 empty] | 0 | - |  | 23 |  | 0 | Common Qualifier Flag | [5 empty] | 0 | - |  | 23
[7 empty] | - |  |  |  | 0 | Code | Description | [5 empty] | -
[7 empty] | - |  |  |  | 0 | LJ | ID acceptable; Value is Estimate | [5 empty] | -
[7 empty] | - |  |  |  | 0 | LL | ID; Value Biased Low | [5 empty] | -
[7 empty] | - |  |  |  | 0 | LK | ID; Value Biased High | [5 empty] | -
[7 empty] | - |  |  |  | 0 | QX | Does not Meet QC Criteria | [5 empty] | -
[7 empty] | - |  |  |  | 0 | EH | Estimated; Exceeds Upper Range | [5 empty] | -
[7 empty] | - |  |  |  | 0 | ND | No Value Detected; Non Detect | [5 empty] | -
[7 empty] | - |  |  |  | 0 | MD | Value Less Than MDL (<MDL) | [5 empty] | -
[7 empty] | - |  |  |  | 0 | SQ | Value b/t SQL and MDL (>= MDL and < SQL) | [5 empty] | -
[7 empty] | - |  |  |  | 0 | 5 | Outlier | [5 empty] | -
[7 empty] | - |  |  |  | 0 | V | Validated Value | [5 empty] | -
[7 empty] | - |  |  |  | 0 | [7 empty] | -
[7 empty] | - |  |  |  | 0 | [7 empty] | -
[7 empty] | - |  |  |  | 0 | [7 empty] | -
[7 empty] | - |  |  |  | 0 | [7 empty] | -
[7 empty] | - |  |  |  | 0 | [7 empty] | -
[7 empty] | - |  |  |  | 0 | [7 empty] | -
[7 empty] | - |  |  |  | 0 | [7 empty] | -
[7 empty] | - |  |  |  | 0 | [7 empty] | -
[7 empty] | - |  |  |  | 0 | [7 empty] | -
[7 empty] | - |  |  |  | 0 | [7 empty] | -
[7 empty] | - |  |  |  | 0 | [7 empty] | -
SAMPLE TYPE CHANGES
DATE/TIME RANGE
FILE NAME |  | OLD TYPE | NEW TYPE |  | Date | Hour | - | Date | Hour | DATE APPLIED
-
-
-
-

## Calculations
CVS Results | [6 empty] | LCS Results
| Propane |  |  | Toluene |  |  |  | Propane | Toluene
DATE/TIME | 43204 |  |  | 45202 |  |  | DATE/TIME | 43204 | 45202
2025-08-01 | 2.038 | 1.5074 | 4.2042 | 2.0217 | 1.44 | 5.1933 | 2025-08-02 21:00 | 5.7653 | 5.7664
2025-08-02 02:00 | 2.0697 | 1.5308 | 3.5843 | 2.0934 | 1.491 | 3.399 | 2025-08-09 21:00 | 5.8388 | 5.8886
2025-08-03 | 2.0466 | 1.5138 | 4.0746 | 2.0192 | 1.4382 | 5.1884 | 2025-08-16 21:00 | 5.8913 | 5.9576
2025-08-03 01:00 | 2.0208 | 1.4947 | 4.3579 | 2.0862 | 1.4859 | 3.6892 | 2025-08-23 21:00 | 5.8806 | 5.8239
2025-08-04 02:00 | 2.0812 | 1.5393 | 3.2831 | 2.0964 | 1.4932 | 3.2765 | 2025-08-30 21:00 | 5.7098 | 5.8698
2025-08-05 | 2.0127 | 1.4887 | 4.3778 | 2.0025 | 1.4263 | 5.0177
2025-08-06 02:00 | 2.0991 | 1.5526 | 2.7745 | 2.0419 | 1.4543 | 5.0335
2025-08-07 | 2.0669 | 1.5288 | 3.653 | 2.1236 | 1.5125 | 2.1906
2025-08-08 02:00 | 2.106 | 1.5577 | 2.5736 | 2.0937 | 1.4912 | 3.3868
2025-08-09 | 2.1653 | 1.6016 | 1.0648 | 2.083 | 1.4836 | 3.8152
2025-08-10 02:00 | 2.166 | 1.6021 | 1.0511 | 2.143 | 1.5264 | 1.5226
2025-08-10 03:00 | 2.1803 | 1.6126 | 0.79644 | 2.1921 | 1.5613 | 0.45601
2025-08-11 | 2.1916 | 1.621 | 0.62862 | 2.1886 | 1.5588 | 0.5037 | RTS Results
2025-08-12 07:00 | 2.0515 | 1.5174 | 3.9866 | 2.0473 | 1.4582 | 4.9334 |  | Propane | Toluene
2025-08-13 01:00 | 2.0711 | 1.5319 | 3.5493 | 2.0321 | 1.4474 | 5.1547 | DATE/TIME | 43204 | 45202
2025-08-14 | 1.9836 | 1.4672 | 4.1695 | 2.0625 | 1.469 | 4.5403 | 2025-08-02 22:00 | 1.1503 | 1.4762
2025-08-15 02:00 | 2.0284 | 1.5003 | 4.3082 | 2.0266 | 1.4434 | 5.1869 | 2025-08-09 22:00 | 1.1764 | 1.4911
2025-08-16 | 1.9892 | 1.4713 | 4.2423 | 2.0546 | 1.4634 | 4.7638 | 2025-08-16 22:00 | 1.1543 | 1.5482
2025-08-17 02:00 | 1.9958 | 1.4762 | 4.3088 | 1.996 | 1.4217 | 4.8899 | 2025-08-23 22:00 | 1.1202 | 1.4563
2025-08-17 03:00 | 2.0714 | 1.5321 | 3.5417 | 2.0904 | 1.4889 | 3.5208 | 2025-08-30 22:00 | 1.1285 | 1.4804
2025-08-18 | 2.0559 | 1.5206 | 3.8997 | 2.0501 | 1.4602 | 4.8729
2025-08-19 02:00 | 1.8808 | 1.3911 | 1.5512 | 1.9921 | 1.4189 | 4.7983
2025-08-20 | 2.0349 | 1.5051 | 4.2427 | 2.0307 | 1.4464 | 5.1654
2025-08-21 15:00 | 1.9055 | 1.4094 | 2.2095 | 1.9435 | 1.3843 | 3.0535
2025-08-22 | 1.9747 | 1.4606 | 4.0249 | 1.9485 | 1.3878 | 3.2585 | Target Line
2025-08-23 02:00 | 2.01 | 1.4867 | 4.3768 | 1.9726 | 1.405 | 4.1996 | Y Axis | X Axis
2025-08-24 | 1.9563 | 1.447 | 3.6303 | 1.9544 | 1.392 | 3.499 | 12 | 1
2025-08-24 01:00 | 1.9433 | 1.4374 | 3.2931 | 1.9542 | 1.3919 | 3.4909 | 11 | 1
2025-08-25 02:00 | 1.8668 | 1.3808 | 1.2287 | 1.9195 | 1.3672 | 2.1074 | 10 | 1
2025-08-26 | 1.9112 | 1.4136 | 2.3726 | 1.924 | 1.3704 | 2.276 | 9 | 1
2025-08-27 02:00 | 1.9255 | 1.4242 | 2.7881 | 1.919 | 1.3668 | 2.089 | 8 | 1
2025-08-28 | 1.925 | 1.4238 | 2.7736 | 1.8934 | 1.3486 | 1.2603 | 7 | 1
2025-08-29 02:00 | 1.9409 | 1.4356 | 3.2272 | 1.8916 | 1.3473 | 1.2112 | 6 | 1
2025-08-30 | 1.987 | 1.4697 | 4.2154 | 2.004 | 1.4274 | 5.0425 | 5 | 1
2025-08-31 02:00 | 1.8283 | 1.3523 | 0.57302 | 1.9419 | 1.3831 | 2.988 | 4 | 1
2025-08-31 03:00 | 1.8531 | 1.3706 | 0.95597 | 1.9818 | 1.4115 | 4.5082 | 3 | 1
|  | 0 | 5.9437e-106 |  | 0 | 1.3987e-150 | 2 | 1
|  | 0 | 5.9437e-106 |  | 0 | 1.3987e-150 | 1 | 1
|  | 0 | 5.9437e-106 |  | 0 | 1.3987e-150 | 0 | 1
|  | 0 | 5.9437e-106 |  | 0 | 1.3987e-150
|  | 0 | 5.9437e-106 |  | 0 | 1.3987e-150
|  | 0 | 5.9437e-106 |  | 0 | 1.3987e-150
*See Validation SOP for calculations.
Ambient Air Spike Evaluation
Ambient Air Hours (Pre and post 1 hour if possible)
| Ethane | Ethylene | Propane | Propylene | Iso-butane | N-butane | Acetylene | Trans-2-butene | 1-butene | Cis-2-butene | Cyclopentane | Iso-pentane | N-pentane | 1,3-butadiene | Trans-2-pentene | 1-pentene | Cis-2-pentene | 2,2-dimethylbutane | 2,3-dimethylbutane | 2-methylpentane | 3-methylpentane | Isoprene | 2-methyl-1-pentene | 1-hexene | N-hexane | Methylcyclopentane | 2,4-dimethylpentane | Benzene | Cyclohexane | 2-methylhexane | 2,3-dimethylpentane | 3-methylhexane | 2,2,4-trimethylpentane | N-heptane | Methylcyclohexane | 2,3,4-trimethylpentane | Toluene | 2-methylheptane | 3-methylheptane | N-octane | Ethylbenzene | M&p-xylene | Styrene | O-xylene | N-nonane | Iso-propylbenzene | Alpha-pinene | N-propylbenzene | M-ethyltoluene | P-ethyltoluene | 1,3,5-tri-m-benzene | O-ethyltoluene | Beta-pinene | 1,2,4-tri-m-benzene | N-decane | 1,2,3-tri-m-benzene | M-diethylbenzene | P-diethylbenzene | N-undecane | N-dodecane | TNMHC | TNMTC
2025-08-01 17:00 | 2.3846 | 0.6905 | 2.0376 | 0.5234 | 2.3785 | 2.6848 | 0.1106 | 0 | 0.0905 | 0.0653 | 0.1268 | 4.2737 | 1.8697 | 0 | 0 | 0 | 0 | 0.0907 | 0.3939 | 0.9034 | 0.5139 | 0.1337 | 0 | 0 | 0.8082 | 0.2947 | 0.3566 | 0.3318 | 0.1823 | 0.2374 | 0.2635 | 0.2259 | 0.5335 | 0.3427 | 0.2655 | 0.1307 | 0.9572 | 0.1292 | 0.1104 | 0.1998 | 0.0922 | 0.4548 | 0 | 0.1484 | 0.2149 | 0 | 0.1342 | 0 | 0.0912 | 0.0302 | 0 | 0 | 0.0414 | 0.0981 | 0.1567 | 0.0343 | 0 | 0 | 0.0785 | 0.0491 | 26.265 | 26.265
2025-08-01 18:00 | 2.3517 | 0.8004 | 2.5327 | 0.6174 | 2.8509 | 5.4678 | 0.09 | 0.0479 | 0.1113 | 0.0658 | 0.1249 | 10.5891 | 2.86 | 0.0383 | 0 | 0.0599 | 0 | 0.1081 | 0.3497 | 0.7758 | 0.4148 | 0.2647 | 0 | 0 | 0.6386 | 0.247 | 0.2017 | 0.3228 | 0.1084 | 0.2099 | 0.1774 | 0.1716 | 0.536 | 0.3875 | 0.2226 | 0.1313 | 0.9997 | 0.1509 | 0.0914 | 0.1647 | 0.0976 | 0.4739 | 0 | 0.1691 | 0.1226 | 0 | 0.1039 | 0.0175 | 0.1098 | 0 | 0 | 0 | 0.0485 | 0.115 | 0.1009 | 0.0391 | 0 | 0 | 0.0655 | 0.0424 | 36.7885 | 36.7885
2025-08-01 22:00 | 1.6934 | 0.2563 | 0.7258 | 0.1335 | 0.2492 | 0.6184 | 0.0692 | 0 | 0.0371 | 0.042 | 0.0361 | 0.8065 | 0.4253 | 0.0274 | 0 | 0 | 0 | 0 | 0.0696 | 0.232 | 0.1281 | 0.6407 | 0 | 0 | 0.146 | 0.0195 | 0.0827 | 0.1647 | 0 | 0.0593 | 0.0884 | 0.049 | 0.1572 | 0.0864 | 0.0568 | 0.0366 | 0.5432 | 0.0413 | 0.0351 | 0.0394 | 0.0717 | 0.3626 | 0 | 0.1053 | 0.0384 | 0 | 0.1347 | 0.0946 | 0.1547 | 0.0632 | 0.069 | 0 | 0.0692 | 0.121 | 0.0857 | 0.0478 | 0 | 0.1114 | 0.0648 | 0.1213 | 9.5116 | 9.5116
Average | 2.1432 | 0.5824 | 1.7654 | 0.42477 | 1.8262 | 2.9237 | 0.089933 | 0.015967 | 0.079633 | 0.0577 | 0.095933 | 5.2231 | 1.7183 | 0.0219 | 0 | 0.019967 | 0 | 0.066267 | 0.27107 | 0.63707 | 0.35227 | 0.34637 | 0 | 0 | 0.53093 | 0.18707 | 0.21367 | 0.2731 | 0.0969 | 0.16887 | 0.17643 | 0.14883 | 0.4089 | 0.2722 | 0.18163 | 0.099533 | 0.83337 | 0.10713 | 0.078967 | 0.13463 | 0.087167 | 0.43043 | 0 | 0.14093 | 0.1253 | 0 | 0.12427 | 0.037367 | 0.11857 | 0.031133 | 0.023 | 0 | 0.053033 | 0.11137 | 0.11443 | 0.0404 | 0 | 0.037133 | 0.0696 | 0.070933 | 24.1884 | 24.1884
*See Validation SOP for calculations.
Ambient Air Spiked Hour
| Ethane | Ethylene | Propane | Propylene | Iso-butane | N-butane | Acetylene | Trans-2-butene | 1-butene | Cis-2-butene | Cyclopentane | Iso-pentane | N-pentane | 1,3-butadiene | Trans-2-pentene | 1-pentene | Cis-2-pentene | 2,2-dimethylbutane | 2,3-dimethylbutane | 2-methylpentane | 3-methylpentane | Isoprene | 2-methyl-1-pentene | 1-hexene | N-hexane | Methylcyclopentane | 2,4-dimethylpentane | Benzene | Cyclohexane | 2-methylhexane | 2,3-dimethylpentane | 3-methylhexane | 2,2,4-trimethylpentane | N-heptane | Methylcyclohexane | 2,3,4-trimethylpentane | Toluene | 2-methylheptane | 3-methylheptane | N-octane | Ethylbenzene | M&p-xylene | Styrene | O-xylene | N-nonane | Iso-propylbenzene | Alpha-pinene | N-propylbenzene | M-ethyltoluene | P-ethyltoluene | 1,3,5-tri-m-benzene | O-ethyltoluene | Beta-pinene | 1,2,4-tri-m-benzene | N-decane | 1,2,3-tri-m-benzene | M-diethylbenzene | P-diethylbenzene | N-undecane | N-dodecane | TNMHC | TNMTC
2025-08-01 19:00 | 9.5148 | 0 | 9.1141 | 0.2779 | 0.1054 | 9.0463 | 3.9522 | 0 | 0 | 0 | 0 | 0.0471 | 8.9183 | 9.2723 | 0 | 0 | 0 | 0 | 0.0643 | 9.0839 | 0.0392 | 0 | 0 | 8.4884 | 8.5472 | 0 | 0 | 9.1283 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 9.2058 | 0.0202 | 0 | 0 | 0.0197 | 9.4817 | 0.0284 | 0.0932 | 0 | 0 | 0.0591 | 9.1308 | 0.0291 | 0 | 0.0244 | 0.0359 | 0.0169 | 7.8782 | 0.0314 | 0.1251 | 0.066 | 6.3307 | 0 | 0.0373 | 128.2136 | 128.2136
Abs. Recovery | 7.3716 | -0.5824 | 7.3487 | -0.14687 | -1.7208 | 6.1226 | 3.8623 | -0.015967 | -0.079633 | -0.0577 | -0.095933 | -5.176 | 7.2 | 9.2504 | 0 | -0.019967 | 0 | -0.066267 | -0.20677 | 8.4468 | -0.31307 | -0.34637 | 0 | 8.4884 | 8.0163 | -0.18707 | -0.21367 | 8.8552 | -0.0969 | -0.16887 | -0.17643 | -0.14883 | -0.4089 | -0.2722 | -0.18163 | -0.099533 | 8.3724 | -0.086933 | -0.078967 | -0.13463 | -0.067467 | 9.0513 | 0.0284 | -0.047733 | -0.1253 | 0 | -0.065167 | 9.0934 | -0.089467 | -0.031133 | 0.0014 | 0.0359 | -0.036133 | 7.7668 | -0.083033 | 0.0847 | 0.066 | 6.2936 | -0.0696 | -0.033633 | 104.0252 | 104.0252
Target | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5
% Recovery | 147.4% | -11.65% | 147% | -2.937% | -34.42% | 122.5% | 77.25% | -0.3193% | -1.593% | -1.154% | -1.919% | -103.5% | 144% | 185% | 0% | -0.3993% | 0% | -1.325% | -4.135% | 168.9% | -6.261% | -6.927% | 0% | 169.8% | 160.3% | -3.741% | -4.273% | 177.1% | -1.938% | -3.377% | -3.529% | -2.977% | -8.178% | -5.444% | -3.633% | -1.991% | 167.4% | -1.739% | -1.579% | -2.693% | -1.349% | 181% | 0.568% | -0.9547% | -2.506% | 0% | -1.303% | 181.9% | -1.789% | -0.6227% | 0.028% | 0.718% | -0.7227% | 155.3% | -1.661% | 1.694% | 1.32% | 125.9% | -1.392% | -0.6727% | 2081% | 2081%
*Add explanation of additional calculations in Case Narrative.
Miscellaneous Calculations and Verifications